规 格 承 认 书

| 客户名称 |  | 红外线发射管实图 型号：CHQ660/905 |
| --- | --- | --- |
| 业务联系人 |  |  |
| 联系电话 |  |  |

| APPROVED SIGNATURES |  |  | APPROVED SIGNATURES |  |  |
| --- | --- | --- | --- | --- | --- |
| 承办 | 审核 | 核准 | 检测 | 审核 | 核准 |
| PREPARED BY | CHECKED BY | APPROVED BY | TESTED BY | CHECKED BY | APPROVED BY |

型号:CHQ660/905

1、Features (特征)

●Lead frame molded packages

引脚框架成型封装

●Two and three or four leads designs

2个或3个或4个引脚设计

●Bi-wavelengths or triple wavelengths LEs

两倍或三倍波长LED

●Matching detector  response

匹配探测器响应

2、Applications (应用领域)

●SPO2    血氧饱和度	● Medical instrumentation 医疗器械

●Blood anaiysis   血液分析  ●Radiometric instruments 辐射仪

| Name 名称 | Model型号 | RED | IR | Package |
| --- | --- | --- | --- | --- |
| Multiwavelength transmitter 多波长发射器 | CHQ660/905 | 660nm | 905nm | 2-pin,COB |

型号:CHQ660/905

| Paramete 参数 | Symbol 符合 | Max 最大 | Unit 单位 | Note 标准 |
| --- | --- | --- | --- | --- |
| Power Dissipation 功率消耗 | Pd | 60 | Mw | … |
| Forward Current 正向电流 | Lf | 20 | Ma | … |
| Peak Forward Current 峰值正向电流 | IFP | 100 | Ma | 1/10 duty cycle,0.1ms pulse width 1/10周期, 0.1ms脉宽 |
| Reverse Voltage 反向电流 | VR | 5 | v | … |
| Operating Temperature 工作温度 | Topr | -25～+85 | ℃ | … |
| Storage Temperature 储藏温度 | Tstg | -40～+100 | ℃ | … |
| Soldering Temperature 焊接温度 | Tsol | 260 | ℃ | 260℃ for 3 Secconds |

3、Absolute Maximum Ratings (ta=25°)在25°环境下的绝对最大额定参

4、Electrical/Optical Characteristics Ta=25℃(在25°环境下的电性/光学特性)

| Parameter 参数 | Symbol 符号 | Min. |  | Typ. |  | Max. |  | Units | Test Conditions 测试条件 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | IRED | RED | IRED | RED | IRED | RED |  |  |
| Forward Voltage 正向电压 | VF | –– | 1.8 | 1.35 | –– | 1.7 | 2.4 | V | IF=20mA |
| Reverse Current 反向电压 | IR | –– | –– | –– | –– | 10 | 10 | Ua | VR=5V |
| Radiant Power 辐射功率 | Po | 1.5 | 9 | –– | –– | 2.5 | 11 | Mw | IF=20mA |
| Peak Waveiength 峰值波长 | λp |  |  | 905 | 660 | –– | –– | nm | IF=20mA |
| Spectral Line Half-width半值波长 | Δλ | –– | –– | 50 | 15 | –– | –– | nm | IF=20mA |

型号:CHQ660/905

5、Dimension  (尺寸)

Notes: 1.All dimensions are in millimeters（尺寸都是以毫米为单位）

2. Tolerances unless dimensions ±0.1mm （未设误差值在±0.1mm）

6、Storage and Soldering Condition

1. Do not open moisture proofs bag before the products are ready to use

在产品准备使用前不要打开防潮袋

2. Before opening the package, the LEDs should be kept at 30℃ or less and 90% RH or les

在打开封装之前产品应该保持在30℃或低于30℃和湿温相对度在90%或小于90%的环境中

3. The LEDs should be used within a year.

LED使用期限为1年

4. After opening the package, the LEDs should be kept at 30℃ or less and 70% RH or less.

在打开封装之后，LED应处于30℃或低于30℃和湿温相对度在70%或低于70%的环境中。

5. The LEDs should be used within 168 hours (7 days) after opening the package.

LED应该在打开封装后的168（7天）内使用

6.When soldering, do not put stress on the LEDs during heating.

在焊接加热时不要对LED上施加压力

7. If the moisture adsorbent material (silica gel) has fabled away or the LEDs have

如果湿气吸附剂材料脱落或者LED超过使用时间

exceeded the storage time, baking treatment should be performed using the following

conditions. Baking treatment: 60±5℃ for 24 hours

应在以下条件进行烘烤处理，烘烤处理在60±5°下进行24小时烘烤。

8. After soldering, do not warp the circuit board

在焊接后，不要扭曲电路板。

9. Each terminal is to go to the tip of soldering iron temperature less than 260℃ for 5

每个接口都是在低于焊铁容量25W时，波封焊温度小于260℃，焊接时间持续5秒

seconds within once in less than the soldering iron capacity 25W. Leave tow seconds and

多端口焊接每隔2秒焊接另端口

more intervals, and do soldering of each terminal. Be careful because the damage of the

product is often started at the time of the hand solder.

请小心，因为产品的损坏通常是在手工焊接时容易产生。